## Slide 1
Assessment & Enhancment of the
CSSD 2019 and Beyond

## Slide 2
SPONSORED BY ONESOURCE
Connecticut Central
Service Association
Service from 1994 through today
President 2008 09 10 11 12 13 DONE!!
Former CSSD Tech Supervisor and
Past Manager Middlesex Hospital
AAMI STWG40 (ST79) voting member
IAHCSMM Ortho council past member
IAHCSMM Executive Board 2011-12 2012-2013
IAHCSMM President Elect 2013-2015
IAHCSMM President 2015-2016
IAHCSMM immediate Past President 2016
IAHCSMM Past president
Director Speakers Bureau Oneource
David Jagrosse Consulting L.L.C.President.
Consulting to medical vendors, Speaking/Education and
AAMI Based Audits for medical facilities
203-465-8502
David Jagrosse Consulting LLC CRCST CHL

## Slide 3: SEAVEY HEALTHCARE CONSULTING® STERILE PROCESSING SURGICAL SERVICES . 303-467-0868 office/fax www.seaveyhealthcareconsulting.com Established in 2003 Sterile Processing Best Practices Audit Check Sheet
Hail to my “work mom”
Rose Seavey. Retired!
Expert,Mentor,Guide, Friend
Rose Seavey, MBA, BS, RN, CNOR, CRCST, CSPDT, is president and CEO of Seavey Healthcare Consulting, Inc. in Arvada, CO, having previously served as the director of the sterile processing department at The Children’s Hospital of Denver. She is a member of numerous AAMI working groups focused on developing sterilization and reprocessing standards and is the author of Sterile Processing in Healthcare Facilities: Preparing for Accreditation Surveys.

## Slide 4
Preparing for Accreditation Surveys, 3rd Edition
This book provides valuable tools for preparing for accreditation surveys and maintaining compliance with accreditation requirements as they relate to sterile processing. New to this third edition are best-practice audit tools for immediate-use steam sterilization (IUSS) and high-level disinfection (HLD), and an instrument integrity checklist. It has also been updated to reflect the 2017 revision of ANSI/AAMI ST79 and current TJC accreditation standards related to sterile processing. TJC's new SAFER matrix scoring methodology is also discussed.
BUY THIS BOOK!!!!! AAMI.ORG

## Slide 5: The efficacy of any sterilization process depends on four phases: 1. A consistent system for lowering and limiting bioburden before sterilization, 2. Properly preparing items for sterilization, 3. Selecting the appropriate sterilization parameters, and 4. Establishing and implementing controls to maintain the sterility of sterilized items until they are used.

## Slide 6
The delivery of sterile healthcare products for use in patient care depends not only on the efficacy of the sterilization process itself but also on the following five factors:
*efficient facility design,
*proper training of personnel,
*good infection prevention and control practices designed to prevent health-care-associated infections,
*effective quality controls and process improvement systems that encompass all aspects of device reprocessing from point of use through sterilization to reuse, and
*appropriate documentation and reporting practices that enable traceability to the patient.

## Slide 7
PERFORMING AUDITS
Critical aspects of infection prevention include proper reprocessing procedures and sterility maintenance. Facilities should perform routine audits for all reprocessing and sterilization areas. These audits should be conducted by the reprocessing manager and the IP.
The following is a list of important issues to monitor based on ANSI/AAMI ST79:2017 Comprehensive guide to steam sterilization and sterility assurance in health care facilites. Some monitoring issues are based on AORN guidelines.

## Slide 9: MAKE SURE FACILITIES IS FOLLOWING NEW ASHRAE STANDARD 170 PER AAMI ST79 2017

## Slide 10
Ball in the wall monitors. Hundreds of dollars plus install cost is best practice

## Slide 11
Any door should have a window for staff safety

## Slide 12: Ensure plugs in decontam or wet areas are GFI protected

## Slide 13
STEAM QUALITY

## Slide 14
WETPACKS

## Slide 17
David Jagrosse Consulting LLC can have review done by Dr Jonathan Wilder

## Slide 18
Resume building!
IFUs!

## Slide 19: DON’T BECOME A TRAINING FACTORY FOR THE HOSPITAL UP THE STREET!!!

## Slide 20
*Repair Program
*Cleaning Verification
*New Staff training
*Schedule and duties

## Slide 22
WE WILL GET BACK TO THIS LATER AND PUT IT ALL TOGETHER

## Slide 23
REPAIRS AND IFUS-UPDATED ACCESSIBLE AND FOLLOWED

## Slide 25: Soiled Instruments from OR

## Slide 26
AAMI ST:79 2017 Standards - Point of Use Cleaning
6.3.1 - Throughout the procedure, instruments should be wiped, as needed, with sterile moistened surgical sponges to remove gross soil. Cannulated instruments should be irrigated with sterile water, as needed, without creating aerosols.
6.3.2 - Gross soil should be removed as soon as possible
6.3.3 - All instruments opened in the procedure room should be considered contaminated
6.3.4 - Instruments should be opened and disassembled
When handling contaminated items personnel should wear appropriate PPE (4.5), remove soil by a method that does not promote cross-contamination (avoid splashing areas near sinks).
Prior to transport, instruments should be prepared in such a way as to prevent organic soils from drying (moistened towel, humipak, or pretreatment product)

## Slide 27
AAMI Standards - Transport of Soiled Instruments
OSHA requires that:
all containers, devices, or carts used for containing contaminated items be marked with a biohazard label, a red bag, or other means of identifying contaminated contents; and
puncture-resistant, leak proof on the sides and bottom, closable, and labeled containers must be used or devices with edges or points capable of penetrating container or skin

## Slide 28
Observed instruments returned to decontam in closed position.

## Slide 29
DeMayo positioner returned without gross soil removed. This should be disposed of in OR.

## Slide 30
Camera returned to decontam underneath other instruments
There goes $5K and some turn overs for tomorrow!!

## Slide 31
DECONTAM AND CLEANING VERIFICATION PROCESS

## Slide 32: Cleaning Verification Program
1. People
  Training. Lead tech in charge of program
2. Process
  Procedure. Checklists, Maintenance of equipment.
3. Mechanical
  Validate the efficacy of equipment and chemistries. In wash Tests. Post wash tests. Critical Water

## Slide 34
All Clean Test logs not completely filled out.
Interpretation chart needs to be available to determine if failures are impingement or chemistry related. (could be available, just not observed)

## Slide 35: Make sure Cart washer maintenance program includes cleaning the holding tank

## Slide 36
Cleaning verification program will include washer maintenance
(spray arm check and clean, descale, etc)

## Slide 37: Review Cycle profiles for washers from MFG as part of review and ensure staff are using appropriate cycles

## Slide 38
Instruments are sent through the wash in closed position
Unsafe expectation for SPD employees to dig in soiled instruments to separate and open
*Picture is from set in assembly area waiting to be assembled

## Slide 39
Linting pipe cleaners being used to brush lumens and to hold instruments open. These shed fibers.
Other non-linting options available
Disposable ex - Healthmark
Key Surgical has reusable brushes (follow IFU for when to throw away and how often to clean brushes)

## Slide 40
Key Surgical or other brush organizer should be on the wall near the hand washing sink.
Appropriate brushes need to be available for suctions (7fr, 9fr, etc)

## Slide 41
Where is clean brush storage? Assembly?
Clean storage area in decontam to replenish when decontam cheklists are done
Accessible to decontam workers

## Slide 42: General clean up. Remove items that are not for immediate usage

## Slide 43
Many cleaning products across the spectrum. Various PH on product. Consider opportunity to switch to one product for washers sinks and cart washer and certified wash cycles. “Milk” bottles on assembly
Use a true enzyme for more effective breakdown

## Slide 44
Cleaning activities duplicated in each bay of the sinks
Bays of sinks should follow a linear process flow
Need to review “Suds” product
No critical water (RO, deionized) available for final rinse
No pressurized medical air available
Wash
Remove gross soil
Brush lumens
Rinse
RO Water Rinse
Soiled >>> Clean

## Slide 45
Items that can go through washer should go through ultrasonic.
Cannulated items should go through sonic with irrigator flush ports
Available from Steris

## Slide 46
No hand washing sink available in decontam.
Because of the paper towel dispenser, it looks like hand washing would take place in decontam. This practice would not follow AAMI guidelines.

## Slide 48
Need protective covers ion scopes to prevent further damage
Assign down sets to Service Leader repair tech

## Slide 49
Assembly staging is area of greatest opportunity
Typically unloaded and organized by specialty
Searching for exceptions and priorities??
Someone has to decide what to do next
When is this set going to get done???
WE SHALL DISCUSS LATER

## Slide 50
Disposable tray liners are better than using towels!

## Slide 51
GO TO HEALTHMARK AND USE CREPE PAPER AUTOCLAVE LINER
*Only to be used under wrapped trays, NOT peel pouches or rigid containers

## Slide 52
No storage under or directly on washer racks on cleanside. Even if temporary

## Slide 53
ALL insulated instruments must be inspected with an insulation tester. New AAMI amendments coming

## Slide 54
Check IFU

## Slide 55
CAN ALMOST BE SEPARATE AUDIT

## Slide 56: Make sure staff are 100% following exposure cycle times per IFU regardless if it is in a Onetray. Dry time is o (zero) but has nothing to do with exposure. This device is NOT Recalled. Use per IFU. Contact David for more guidance davidjagrosse@gmail.com

## Slide 58: AAMI ST:79 2017 Standards and Wrap IFU
AAMI ST70 2017 11.1.1
Positioned so that packaging is not crushed, bent, compressed, or punctured
Wrap manufacturer’s IFU should be consulted for guidance on stacking trays
Open shelving may be used, but requires special attention to traffic control, area ventilation, and environmental services
The bottom shelf of storage carts or shelving should be solid

## Slide 59
Review IFU for tray belt. Currently used in CSS for sterilizer cart. Intended purpose is for sterile storage and OR Mishandling
For Storage and external handling
For sterilizer cart

## Slide 60
In-servicing on wrapping
Wrap was being done inconsistently and incorrectly
Get Halyard,Cardinal or Medline on-site to do education

## Slide 61
Get window coverings for windows in storage area
UV can degrade packaging and contents

## Slide 62
Need different storage solution for long peel pouches

## Slide 63
Check IFU’s to verify these practices

## Slide 64
Wrapped sets should not be stacked on top of rigid containers

## Slide 65
Use of Halyard blue trays
stack tray on tray for better protection

## Slide 66
Shelving also needs to be reviewed in SPD

## Slide 69
Consider Quality management for future. AAMI ST90:2017

## Slide 71
Assesment/Audit internal or external person
AUDITS

## Slide 72
Hospitals have a capital budget process for approving new devices. Part of that process is a form that needs to be filled out. It is within that form that we have the opportunity to identify devices that require high-level disinfection or sterilization. Change the request form to ask this question and if the box is checked off as yes it should go to the cssd manager to ensure that we have the ability to reprocess the device.

## Slide 73
Policies and documented procedures are the most powerfull tool you have for accuracy and consistency within the system. They are also essential when DPH or JCAHO inspections occur. Make sure your polices are current and based on best practice,MFGR IFU and current guidelines.

## Slide 74: Process Improvement Loaner Process Forecasting → Processing → Staging Assembly Productivity T-Processing Case Processing Traffic Control Set up processes now to prevent transfer of current habits created by logistical constraints to new space

## Slide 77
CURRENT STATE:
“ISLANDS of Processing”
Washer output
Washer output
Washer output
ASSEMBLY ISLAND
ASSEMBLY ISLAND
ASSEMBLY ISLAND
ASSEMBLY ISLAND
WRAP
WRAP
WRAP
WRAP
Duplication of peel pouching,wrap insulation testing etc.
No team work
Cherry Pick sets
Poor work flow
No Quality checks by design.
CONTAINER CART

## Slide 78
Eliminate holding racks by going to “T” processing and new prioritization system and processing by case with tags in decontam.

## Slide 79
Cases
1 and 2
Cases
4 and 6
Case 3
Case 5

## Slide 84
ALL insulated instruments must be inspected with an insulation tester. New AAMI amendments coming

## Slide 86
FUTURE STATE

## Slide 87
Make sure your AAMI ST79 is latest version and accessible. 2017

## Slide 88
Involvement with state IAHCSMM chapter (CCSA). Staff attend local education conferences for CEU. Hospital reimburse cost of attendance.
Leadership in dept to attend annual IAHCSMM conference for education and networking. Also possible involvement with AAMI ST:79 workgroup
Start Library based out of office of educational books for staff to read and develop.
Free magazines for education/CEU’s. Infection Controol today,Healthcare Purchasing News, IAHCSMM magazine.

## Slide 90: Join the Fight
"Those who expect to reap the blessings of freedom must, like men, undergo the fatigue of supporting it." Thomas Paine

## Slide 91: Change will not come about through expectation –Dave Jagrosse

## Slide 92
THANK YOU!